| Method comparison | Comparison of ICP, IC and spectrophotometric determination in coloured soil solutions | SO4 |
| --- | --- | --- |

# 1.	Compared methods:

- Ion exchange chromatography without chemical suppression; eluent: phthalic acid with tris(hydroxymethyl)-aminomethane; column: Shimpack IC-A1S; pre-column: Shimpack IC-GA1; conductivity detector (König, Fortmann 1996a).

- Stotal determination with ICP, correction for organic sulphur: determination of DOC and subtraction of Corg [mg/l]/130 from total sulphur content; ICP with axial plasma (Polyscan 61E, Thermo Instruments); wavelength: 182.040 nm; cross-flow nebulizer; without internal standard; no interelement correction (König, Fortmann 1996b).

- CFA spectrophotometric determination of SO4 with Ba-methylthymol blue reaction after separation of interfering cations with a cation exchange column; elimination of the interference of humic acids by measuring the blind reaction with Ca-methylthymol blue in a second CFA channel  (König, Fortmann 1996c).

# 2.	Procedure:

The soil solutions were filtered with membrane filters (0.45 µm). About 500 coloured soil solutions were measured with the different SO4 methods.

# 3.	Results:

Using a pre-column reduces the interference of the humic acids in IC. The IC method was therefore used as the reference method. After correction for organic sulphur the ICP method gave comparable results. The spectrophotometric method wasn’t able to correct for all the interferences of the humic acids and their different reactions, but gave much better results than without correction with the blind second channel. Therefore, the CFA spectrophotometric method is not useful for low SO4 concentrations.

(See Fig. 1)

# 4.	Literature:

- König, Fortmann 1996a: Probenvorbereitungs-, Untersuchungs- und Elementbestimmungsmethoden des Umweltanalytik-Labors der Niedersäch-sischen Forstlichen Versuchsanstalt und des Zentrallabor II des Forschungs-zentrums Waldökosysteme, Teil 2: Elementbestimmungsmethoden N-Z, Berichte d. Forschungszentrums Waldökosysteme , Band B 47, 1996; Method SSO4IC1.1

- König, Fortmann 1996b: Probenvorbereitungs-, Untersuchungs- und Elementbestimmungsmethoden des Umweltanalytik-Labors der Niedersäch-sischen Forstlichen Versuchsanstalt und des Zentrallabor II des Forschungs-zentrums Waldökosysteme, Teil 2: Elementbestimmungsmethoden N-Z, Berichte d. Forschungszentrums Waldökosysteme , Band B 47, 1996; Method SSgesICP1.2 and SSO4ICP1.1

- König, Fortmann 1996c: Probenvorbereitungs-, Untersuchungs- und Elementbestimmungsmethoden des Umweltanalytik-Labors der Niedersäch-sischen Forstlichen Versuchsanstalt und des Zentrallabor II des Forschungs-zentrums Waldökosysteme, Teil 2: Elementbestimmungsmethoden N-Z, Berichte d. Forschungszentrums Waldökosysteme , Band B 47, 1996; Method SSO4CFC1.2

# 5.	Executive laboratory:

Niedersächsische Forstliche Versuchsanstalt

Nils König

Grätzelstrasse 2

D 37079 Göttingen

Germany

E-mail: nils.koenig@nfv.gwdg.de

6. Tables and graphs, appendix:

Figure 1: