| (B) | EPD info sheets: General information (C) | (D) |
| --- | --- | --- |

1. Why EPD info sheets:

In all good laboratories, much experience is gained during daily work. Different materials are tested, analytical methods are compared, new methods are tested or possibilities for storage and preparation of samples are examined. Often this collected experience is only available to a few persons working in the laboratory, because there is no time to write a special publication about such tests or comparisons.

With the EPD info sheets, we try to make this experience available for all members of the Expert Panel on Deposition and the authorized laboratories of the member countries.

We want to motivate you to write down such short communications about your practical experience. We have therefore developed 5 different Word files that can be filled out in a short and easy way.

2. The different Word files for EPD info sheets:

We have developed 5 different Word files for

- method comparisons (EPDMC)

- new methods (EPDNM)

- material tests (EPDMT)

- sample pretreatment and storage (EPDSP)

- technical information (EPDTI)

3. How to use the Word files for EPD info sheets:

First you have to decide which Word file you want to use to write down your results or experience.

Then you have to fill in two of the six boxes in the two heading lines:  Boxes C and D. Box C should be filled in with the title of the analytical information and Box D with the ion (MC and NM files) or the kind of sample (SP, TI and MT files). Box B is already filled in with the type of information (method comparison, new method, material test, sample pretreatment and storage or technical information) and Box A will be filled in by us with a reference number for the info sheet. Please don’t change the format of the two heading lines!

Then you have to write your text under the 6 or 7 predetermined section headings of the Word file. Please don’t change the section headings!

The contact details are very important, because in this way it is possible to ask for more and detailed information.

If you need more than one side for your info sheet, there is no problem: if you write a longer text the Word file will generate the next side automatically.

4. Examples of EPD info sheets:

Two examples of EPD info sheets already exist:

MC1: Comparison of ICP, IC and spectrophotometric determination in coloured soil solutions

MT1: Leaching of PE bottles

5. Where to send the filled in Word files:

Please send the filled in Word file by e-mail to the following e-mail address:

nils.koenig@nfv.gwdg.de

2 or 3 members of the working group will read the files that you have sent to eliminate unclarity. The collected EPD info sheets will be sent out every 2 months to all expert panel members by e-mail.

6. Questions/problems:

If there are any questions or problems with the Word files don’t hesitate to ask me!

Contact:

Nils König

Niedersächsische Forstliche Versuchsanstalt

Grätzelstrasse 2

D 37079 Göttingen

Germany

Tel: 0049-551-69401141

Fax: 0049-551-69401160

E-mail: nils.koenig@nfv.gwdg.de

7. Appendix:

The two above-mentioned examples of EPD info sheets and the five Word files are attached to this text. Please don’t copy the attached files for your info sheets on account of problems with formatting. Instead, for your info sheets please use the extra files:

- EPDMT

- EPDNM

- EPDSP

- EPDMC

- EPDTI

| Method comparison | Comparison of ICP, IC and spectrophotometric determination in coloured soil solutions | SO4 |
| --- | --- | --- |

# 1.	Compared methods:

- Ion exchange chromatography without chemical suppression; eluent: phthalic acid with tris(hydroxymethyl)-aminomethane; column: Shimpack IC-A1S; pre-column: Shimpack IC-GA1; conductivity detector (König, Fortmann 1996a).

- Stotal determination with ICP, correction for organic sulphur: determination of DOC and subtraction of Corg [mg/l]/130 from total sulphur content; ICP with axial plasma (Polyscan 61E, Thermo Instruments); wavelength: 182.040 nm; cross-flow nebulizer; without internal standard; no interelement correction (König, Fortmann 1996b).

- CFA spectrophotometric determination of SO4 with Ba-methylthymol blue reaction after separation of interfering cations with a cation exchange column; elimination of the interference of humic acids by measuring the blind reaction with Ca-methylthymol blue in a second CFA channel  (König, Fortmann 1996c).

# 2.	Procedure:

The soil solutions were filtered with membrane filters (0.45 µm). About 500 coloured soil solutions were measured with the different SO4 methods.

# 3.	Results:

Using a pre-column reduces the interference of the humic acids in IC. The IC method was therefore used as the reference method. After correction for organic sulphur the ICP method gave comparable results. The spectrophotometric method wasn’t able to correct for all the interferences of the humic acids and their different reactions, but gave much better results than without correction with the blind second channel. Therefore, the CFA spectrophotometric method is not useful for low SO4 concentrations.

(See Fig. 1)

# 4.	Literature:

- König, Fortmann 1996a: Probenvorbereitungs-, Untersuchungs- und Elementbestimmungsmethoden des Umweltanalytik-Labors der Niedersäch-sischen Forstlichen Versuchsanstalt und des Zentrallabor II des Forschungs-zentrums Waldökosysteme, Teil 2: Elementbestimmungsmethoden N-Z, Berichte d. Forschungszentrums Waldökosysteme , Band B 47, 1996; Method SSO4IC1.1

- König, Fortmann 1996b: Probenvorbereitungs-, Untersuchungs- und Elementbestimmungsmethoden des Umweltanalytik-Labors der Niedersäch-sischen Forstlichen Versuchsanstalt und des Zentrallabor II des Forschungs-zentrums Waldökosysteme, Teil 2: Elementbestimmungsmethoden N-Z, Berichte d. Forschungszentrums Waldökosysteme , Band B 47, 1996; Method SSgesICP1.2 and SSO4ICP1.1

- König, Fortmann 1996c: Probenvorbereitungs-, Untersuchungs- und Elementbestimmungsmethoden des Umweltanalytik-Labors der Niedersäch-sischen Forstlichen Versuchsanstalt und des Zentrallabor II des Forschungs-zentrums Waldökosysteme, Teil 2: Elementbestimmungsmethoden N-Z, Berichte d. Forschungszentrums Waldökosysteme , Band B 47, 1996; Method SSO4CFC1.2

# 5.	Executive laboratory:

Niedersächsische Forstliche Versuchsanstalt

Nils König

Grätzelstrasse 2

D 37079 Göttingen

Germany

E-mail: nils.koenig@nfv.gwdg.de

6. Tables and graphs, appendix:

Figure 1:

| Material test | Leaching of PE bottles | Water samples and digestion solutions |
| --- | --- | --- |

# Problem / question:

Cu (from plasticisers) seems to leach out from new PE bottles (bottles from different producers).

# Description of the test:

In each case 5 new 100 ml PE bottles were washed in a washing machine with demineralised water and filled with 50 ml HNO3 solutions of different acid concentrations: 0 %, 2.3 %, 4.6 %, 9.2 %. Then the bottles were shaken for 24 hours. These solutions were measured for Cu with graphite furnace AAS. The bottles were washed again in the washing machine and then filled with 2.3 % HNO3 solution. These solutions were also analysed for Cu with graphite furnace AAS.

- Results:

Cu was leached from all bottles in concentrations up to 1.0 µg/l (mean: 0.6 µg/l) after the first treatment with acid. After the second treatment with 2.3 % acid the measured concentrations decreased down to 0.2 µg/l (mean), only for those treated previously with 4.6 % or 9.2 % acid.

- Consequences:

New PE bottles have to be washed with HNO3 of at least 4.6 % before first use!

# Literature:

---

# Executive laboratory:

Niedersächsische Forstliche Versuchsanstalt

Nils König

Grätzelstrasse 2

D 37079 Göttingen

Germany

E-mail: nils.koenig@nfv.gwdg.de

- Appendix:

----

| Material test | (Tested material) | (Kind of samples the material is used for) |
| --- | --- | --- |

# Problem / question:

# Description of the test:

- Results:

- Consequences:

# Literature:

# Executive laboratory:

Institute

Name of contact person

Street

City

Country

E-mail

- Tables and graphs, appendix:

| New method | (Method title) | (Ion) |
| --- | --- | --- |

# Principle of the method:

- Interferences:

# Method description:

- Instruments and accessories:

- Reagents:

- Solutions:

- Calibration / standards:

- Execution:

- Evaluation and report:

# Literature:

# Executive laboratory:

Institute

Name of contact person

Street

City

Country

E-mail

- Appendix:

| Sample pretreatment and storage method | (Title of the test) | (Kind of sample) |
| --- | --- | --- |

# Problem:

# Description of the test:

- Results:

- Consequences:

# Literature:

# Executive laboratory:

Institute

Name of contact person

Street

City

Country

E-mail

- Tables, graphs, appendix:

| Method comparison | (Title) | (Ion) |
| --- | --- | --- |

# 1.	Compared methods:

# 2.	Procedure:

# 3.	Results:

# 4.	Literature:

# 5.	Executive laboratory:

Institute

Name of contact person

Street

City

Country

E-mail

6. Tables and graphs, appendix:

| Technical Information | (Instrument, sampler etc.) | (Kind of samples the instrument is used for) |
| --- | --- | --- |

# 1.	Instrument description:

# 2.	Instructions for use:

# 3.	Problems:

# 4.	Literature:

# 5.	Executive laboratory:

Institute

Name of contact person

Street

City

Country

E-mail

6. Tables and graphs, appendix: